## IOT Device
分類 | 編號 | 測項 | 判定結果 | 目前情況 | 修補建議 | 連結 | 備註 | Report | Report_ch | cvss | score
Static | S01 | Source code scanning | Not Applicable | 無提供 |  |  |  | As the source code is not provided, the source code scanning cannot be performed, and this test item is not within the scope of this assessment. Therefore, this test item is assessed as not applicable. | 由於未提供原始碼，無法執行原始碼掃描，故本測項不在本次檢測範圍內。因此，此測項評估為不適用。 | - | -
Static | S02 | SBOM Vulnerability Scanning | None | 參考SBOM分頁 |  |  |  | Based on the review of the SBOM tab, no abnormalities or vulnerabilities were found. Therefore, this test item is assessed as no risk. | 經檢視SBOM分頁內容，未發現異常或弱點。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
Dynamic | D01 | Sandbox Test | None | 1.操作裝置，進行測量，無發生異常。 |  |  |  | During sandbox testing, the device was operated to perform measurements, and no anomalies occurred. Therefore, this test item is assessed as no risk. | 透過操作裝置進行測量測試，過程中無發生異常。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
Fuzzing | F01 | Fuzzing Test | None | 透過nRF進行藍牙配對及資料寫入 (如超長字串、特殊字元等) 無發生系統當機，無發現異常。 |  |  |  | During testing, Bluetooth pairing and data writing (such as excessively long strings, special characters, etc.) were conducted via nRF. No system crash occurred, and no abnormalities were observed. Therefore, this test item is assessed as no risk. | 透過 nRF 進行藍牙配對及資料寫入（如超長字串、特殊字元等）進行測試，未發生系統當機且無發現異常。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
PT | P01 | Device Vulnerability Scanning | Not Applicable | 此裝置為iot設備，無主機弱掃介面 |  |  |  | As this device is an IoT device without a host vulnerability scanning interface, this test item is not within the scope of this assessment. Therefore, this test item is assessed as not applicable. | 由於此裝置為IoT設備，無主機弱點掃描介面，故本測項不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P02 | UI Vulnerability Scanning | Not Applicable | 介面僅按鈕操作，無web 介面 |  |  |  | The target interface only supports button operations and does not have a web interface, making UI vulnerability scanning unfeasible. Consequently, this test item is not within the scope of this assessment. Therefore, this test item is assessed as not applicable. | 目標設備之介面僅支援按鈕操作，並無 Web 介面，無法執行介面弱點掃描，故本測項不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P03 | Authentication (Remote) | Not Applicable | 無網路介面 |  |  |  | As the target device possesses no network interface, remote authentication mechanisms and login controls are not present. Consequently, this test item is not within the scope of this assessment. Therefore, this test item is assessed as not applicable. | 由於該設備無網路介面，並無遠端身分驗證機制或登入控制功能，故本測項不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P04 | Authentication (Local) | None | 單機使用，不需認證 |  |  |  | Based on verification, the device is used in standalone mode and does not require an authentication mechanism. Therefore, this test item is assessed as no risk. | 經確認，該設備為單機使用，不需身分驗證機制。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
PT | P05 | Network Security Test (Remote) | Not Applicable | 無網路介面 |  |  |  | Since the target device does not have any network interface, remote network security testing cannot be conducted, and this test item is not within the scope of this assessment. Therefore, this test item is assessed as not applicable. | 由於受測設備無網路介面，無法進行遠端網路安全測試，故此測項不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P06 | Network Security Test (Adjacent) | Low | 藍牙側錄： 1.藍牙版本5.1 2.有開啟傳輸加密 3.對應CVE參考"藍牙5.1 CVE list" (高風險) 藍牙由裝置發起配對，為Just Works模式，不必輸入配對資訊(PIN Code / Passkey)，但因為藍牙由裝置按鈕進行人工啟動，非長時間開啟，且傳輸敏感性資料有進行加密，判斷為低風險 |  |  |  | During Bluetooth packet capture analysis, the device was found to use Bluetooth version 5.1 and operate in Just Works pairing mode initiated by the device without requiring pairing information such as a PIN Code or Passkey, exposing it to vulnerabilities referenced in the Bluetooth 5.1 CVE list. However, Bluetooth must be manually activated via a physical button on the device rather than remaining persistently enabled, and sensitive data transmission is protected by encryption. Therefore, this test item is assessed as low risk. | 於藍牙側錄分析中，裝置使用藍牙 5.1 版本並由裝置發起配對，採用無需輸入 PIN 碼或密碼金鑰（Passkey）等配對資訊的 Just Works 模式，存在藍牙 5.1 CVE 清單對應之相關風險。然而，藍牙功能需由裝置按鈕進行人工啟動，非長時間開啟，且傳輸敏感性資料有開啟加密保護。因此，此測項評估為低風險。 | CVSS:3.1/AV:A/AC:L/PR:N/UI:N/S:U/C:L/I:L/A:N | 5.4
PT | P07 | Minimize Network Service | None | 僅提供藍牙服務，參考P06 |  |  |  | Through the evaluation of network services, it was confirmed that the device only provides Bluetooth service (refer to P06), and no unnecessary network services were found. Therefore, this test item is assessed as no risk. | 透過網路服務評估進行確認，裝置僅提供藍牙服務（參考 P06），未發現其他非必要網路服務。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
PT | P08 | Sensitive Data Leakage | Low | 資料僅能透過藍牙傳輸，藍牙傳輸時資料有加密，參考P06 |  |  |  | Sensitive data is transmitted exclusively via Bluetooth, where the underlying Bluetooth 5.1 implementation has associated vulnerabilities and employs Just Works pairing without requiring pairing information, presenting potential risks of sensitive data exposure during data transmission. However, an encryption protection mechanism is applied during Bluetooth transmission, and Bluetooth is not persistently active as it requires manual activation via a device button. Therefore, this test item is assessed as low risk. | 敏感資料僅能透過藍牙進行傳輸，而底層藍牙 5.1 存在相關漏洞且採用不需輸入配對資訊之 Just Works 配對模式，於資料傳輸時存在潛在的敏感資料暴露風險。然而，藍牙傳輸過程中具備資料加密之保護機制，且藍牙非長時間開啟，需透過裝置按鈕人工啟動。因此，此測項評估為低風險。 | CVSS:3.1/AV:A/AC:L/PR:N/UI:N/S:U/C:L/I:N/A:N | 4.3
PT | P09 | API Authentication | Low | 參考P06 |  |  |  | Based on the reference test item P06, the device initiates Bluetooth pairing using the Just Works mode without requiring authentication credentials such as a PIN code or passkey, indicating a lack of robust authentication mechanism during pairing. However, the risk is mitigated because the Bluetooth function is manually activated via a physical button on the device rather than remaining continuously enabled, and sensitive data transmission is encrypted. Therefore, this test item is assessed as low risk. | 依據參考測項 P06，裝置發起的藍牙配對採用 Just Works 模式，無須輸入 PIN 碼或金鑰等配對資訊，身分驗證機制不足。然而，由於藍牙功能需透過裝置上的實體按鈕進行人工啟動，非長時間持續開啟，且敏感性資料傳輸有進行加密，有效降低了風險。因此，此測項評估為低風險。 | CVSS:3.1/AV:A/AC:L/PR:N/UI:N/S:U/C:L/I:L/A:N | 5.4
PT | P10 | API Input Validation | None | 參考F01 |  |  |  | Testing was conducted by performing Bluetooth pairing and data writing (such as excessively long strings and special characters) using nRF, and no system crashes occurred and no anomalies were observed. Therefore, this test item is assessed as no risk. | 透過 nRF 進行藍牙配對及資料寫入（如超長字串、特殊字元等）進行測試，無發生系統當機，無發現異常。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
PT | P11 | UI: Default Password | Not Applicable | 無帳號密碼設定功能 |  |  |  | The target does not feature account and password configuration functions or authentication mechanisms; therefore, this test item is not within the scope of this assessment. Therefore, this test item is assessed as not applicable. | 此目標並無帳號密碼設定功能或相關身分驗證機制，故不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P12 | UI: Password Strength | Not Applicable | 無帳號密碼設定功能 |  |  |  | The target system does not provide an account and password configuration function, meaning there is no authentication mechanism or password policy to evaluate. Consequently, this test item is not within the scope of this assessment. Therefore, this test item is assessed as not applicable. | 受測系統無帳號密碼設定功能，並無相關身分驗證機制或密碼政策可供評估，故本測項不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P13 | UI: Password Management | Not Applicable | 無帳號密碼設定功能 |  |  |  | The target device does not provide any account or password configuration function; therefore, testing for password policy and credential management cannot be performed and is not within the scope of this evaluation. Therefore, this test item is assessed as not applicable. | 受測設備未提供帳號與密碼設定功能，故無法進行密碼政策與帳號憑證管理之檢測，此測項不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P14 | UI: Configuration Management | None | 介面僅能根據韌體的設計選取對應的功能。無法進行額外設定。 |  |  |  | Observation shows that the interface only allows selecting corresponding functions based on the firmware design, and no additional configurations can be performed. Therefore, this test item is assessed as no risk. | 檢測觀察顯示介面僅能根據韌體的設計選取對應的功能，無法進行額外設定。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
PT | P15 | Security Event Logging | Not Applicable | 無相關設計 |  |  |  | As the product has no relevant design for security event logging, this test item is not within the scope of this testing. Therefore, this test item is assessed as not applicable. | 由於本產品無安全事件日誌的相關設計，故此測項不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P16 | Security Event Notification | Not Applicable | 無相關設計 |  |  |  | Since there is no relevant design for security event notification in the product, this test item is not within the scope of this assessment. Therefore, this test item is assessed as not applicable. | 此測項因為產品無安全事件通知之相關設計，故不在本次測試範圍內。因此，此測項評估為不適用。 | - | -
PT | P17 | Application Update | None | 召回 |  |  |  | Based on the verification of the application update mechanism, the current status is recall. Therefore, this test item is assessed as no risk. | 經確認應用程式更新機制，目前狀態為召回。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
PT | P18 | Application Delivery Protection | None | 召回 |  |  |  | Based on the verification of the application delivery protection mechanism, the current status is recall. Therefore, this test item is assessed as no risk. | 經確認應用程式交付保護機制，目前情況為召回。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
PT | P19 | Static Application Analysis | None | 針對 MCU及藍牙韌體進行敏感性字串偵蒐，兩者皆無發現任何敏感性關鍵字，故評估為無風險。 |  |  |  | Static analysis was conducted on the MCU and Bluetooth firmware to search for sensitive strings, and no sensitive keywords were identified in either firmware. Therefore, this test item is assessed as no risk. | 針對 MCU 及藍牙韌體進行敏感性字串偵蒐，兩者皆未發現任何敏感性關鍵字。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0
PT | P20 | Physical Protection | None | 廠商提供 Read Protection防止韌體讀取設定，故評估為無風險。 |  |  |  | Upon verification, the vendor provides Read Protection settings to prevent firmware from being read. Therefore, this test item is assessed as no risk. | 經確認，廠商已提供 Read Protection 設定以防止韌體被讀取。因此，此測項評估為無風險。 | CVSS:3.1/AV:-/AC:-/PR:-/UI:-/S:-/C:N/I:N/A:N | 0

## Security Recommendation
分類 | 最高風險 | 觸發測項 | Recommendation | Recommendation_ch
Static | no risk | -
Dynamic | no risk | -
Fuzzing | no risk | -
PT | low risk | P06, P08, P09 | For adjacent network communications such as Bluetooth, it is recommended to evaluate strengthening the pairing and authentication mechanisms by considering more secure pairing models with mutual authentication or man-in-the-middle protection instead of unauthenticated Just Works pairing, where feasible. In addition, the organization should establish a vulnerability tracking and patch management process for the Bluetooth stack and underlying firmware to remediate known CVEs promptly, while continuing to enforce robust end-to-end cryptographic controls and secure key management for sensitive data transmission to maintain defense-in-depth against unauthorized access and interception. | 針對藍牙等鄰近網路通訊介面，建議評估強化其配對與身分驗證機制，在可行情況下考慮採用具備雙向身分驗證或防中間人攻擊保護之配對模式，取代缺乏身分鑑別的 Just Works 機制；同時應建立藍牙協定堆疊與相關韌體的已知弱點（CVE）追蹤與修補管理流程，並持續在敏感資料傳輸上維持嚴謹的端對端加密與金鑰保護控制，落實縱深防禦以降低非授權連線與通訊攔截的風險。

## SBOM
EBPHKW6-D
Project | package | issue | version | 修正版本 | 建議版本 | Risk | CVE | GHSA | Git | 備註
| CCRL |  | 1.02.00.00 | 未發現可確認之公開 CVE
| CCRL |  | 1.07.00.00 | 未發現可確認之公開 CVE

## 藍牙5.1 CVE list
風險值 | 影響範圍 | CVE ID | 漏洞類型 | 描述 | 影響版本 | 風險 | 藍牙組織揭露
高 | 作業系統 (Android) | CVE-2020-0022 | 遠程代碼執行（RCE） | 攻擊者可通過藍牙執行遠程代碼，控制設備。 | 藍牙4.2和5.x | 敏感資料可被遠程訪問和竊取。
中 | 協議 (BR/EDR/LE) | CVE-2020-10134 | 中間人攻擊 | 攻擊者可以未經授權地完成藍牙設備之間的經驗證配對，從而可能訪問敏感的數據 | 藍牙4.2到5.1 | 利用確認號碼與密碼的不一致，欺騙用戶完成不安全的配對 | ✔
中 | 協議 (BR/EDR) | CVE-2020-10135 | 加密密鑰重設漏洞 | 允許攻擊者重新配對設備，篡改通信加密密鑰。 | 藍牙4.2和5.x | 攻擊者可未經授權訪問設備並竊取數據。
中 | 協議 (BR/EDR/LE) | CVE-2020-26558 | 中間人攻擊 | 攻擊者可以未經授權地完成藍牙設備之間的經驗證配對，從而可能訪問敏感的數據 | 藍牙2.1到5.1 | 攻擊者可以未經授權地完成藍牙設備之間的經驗證配對，從而可能訪問敏感的數據 | ✔
中 | 協議 (LE) | CVE-2020-35473 | 訊息洩漏 | 如果一個 可解析私人地址（Resolvable Private Address, RPA）RPA 與特定的遠端設備建立了關聯，攻擊者可以通過發送主動掃描請求來識別該設備，這被稱為一種基於允許列表（allowlist）的側信道攻擊。 | 藍牙4.2到5.2 | 依賴 RPA 進行隱私保護的裝置，包括智能手機、穿戴設備、物聯網設備等造成隱私洩漏
中 | 協議 (LE) | CVE-2021-31615 | 中間人攻擊 | 在接收端設備的接收窗口內，攻擊者可注入經過設計的數據包。 該數據包可在傳輸端設備開始其數據包傳輸之前成功攔截通信，實現中間人攻擊（MITM）。 | 藍牙4.2到5.2 | 攻擊者可實現完全中間人狀態，攔截並操控未加密通信 | ✔
中 | 協議 (BR/EDR) | CVE-2021-34145 | 拒絕服務（DoS）攻擊 | 特製封包可導致設備崩潰，無法正常運作。 | 藍牙4.2及更新版本 | 設備崩潰，影響藍牙通信。
中 | 協議 (BR/EDR/LE) | CVE-2021-37577 | 中間人攻擊 | 攻擊者在無需身份驗證的情況下，利用公鑰反射及密碼推測，冒充合法設備完成配對。 | 藍牙2.1到5.3 | 1.攻擊者可攔截並解碼配對設備之間的數據通信。 2.攻擊者可冒充合法設備，執行未經授權的操作。 | ✔
中 | 協議 (BR/EDR) | CVE-2022-24695 | 藍牙協議中的設計缺陷 | 攻擊者可通過高效的無線攻擊（over-the-air attack），獲取設備能力（Capabilities），其中可能包含設備擁有者的個人資訊 | 藍牙5.3及早期版本 | 1.設備隱私洩露：攻擊者可提取唯一標識符（如 MAC 地址）和其他設備信息。 2.被動數據收集：攻擊者可通過收集的數據進行設備跟踪或識別。 3.未授權連接風險：攻擊者可能利用 MAC 地址主動嘗試與目標設備建立連接。
高 | 協議 (LE) | CVE-2022-25836 | 中間人攻擊 | 攻擊者可以未經授權地完成藍牙設備之間的經驗證配對，從而可能訪問敏感的數據 | 藍牙4.0到5.3 | 攻擊者可以攔截並操控設備之間的配對過程，偽裝為合法設備。攻擊成功後，攻擊者可能： 1.存取敏感數據。 2.操控 BLE 設備進行未經授權的操作（例如智能家居設備、醫療設備或物聯網設備）。 | ✔
中 | 協議 (BR/EDR/LE) | CVE-2023-24023 | 藍牙會話密鑰派生過程 | 攻擊者可強制使用較短且容易預測的密鑰，從而解密過去和未來的通信內容 | 藍牙4.2到5.4 | 攻擊者可解密或操控設備之間的通信，即使使用了安全連接（Secure Connections）模式。
中 | 協議 (LE) | CVE-2024-33454 | 執行任意程式碼 | 此漏洞存在於 ESP-IDF v5.1 的藍牙堆疊中，特別是在處理 AVDTP（音訊/視訊分發傳輸協定）請求時，未正確驗證輸入資料的大小，導致緩衝區溢位。 | 藍牙5.0到5.1 | 攻擊者傳送特製的 AVDTP 封包，觸發記憶體損毀，導致設備重啟或執行未預期的程式碼。
高 | 作業系統 (Windows) | CVE-2025-27490 | 執行任意程式碼 | 此漏洞源於 Windows 藍牙服務在處理來自藍牙設備的資料時，未正確驗證輸入大小，導致堆積緩衝區溢位。 | 藍牙5.0到5.1 | 攻擊者可利用此漏洞覆寫關鍵記憶體結構，進而執行任意程式碼或提升權限。

## _
Critical
High
Medium
Low
None
Not Applicable
TBD
Testing